## parts
part | type | tag | number | price_per_device_CAD | description | note | provider | link
Pi zero W | off-the-shelf | pi_zero | 1 | 20 | The [Raspberry Pi Zero W](https://www.raspberrypi.com/products/raspberry-pi-zero-w/) microcomputer |  | canakit.com | https://www.canakit.com/raspberry-pi-zero.html?defpid=4307
Pi camera HQ | off-the-shelf | camera | 1 | 70 | The [Raspberry Pi High Quality Camera](https://www.raspberrypi.com/products/raspberry-pi-high-quality-camera/) |  | pishop.ca | https://www.pishop.ca/product/raspberry-pi-hq-camera/
8mm CS lens | off-the-shelf | lens | 1 | 30 | An 8mm Cs mount lens with manually-adjustable aperture and focus and a 50° horizontal field of view. |  | Arducam | https://www.arducam.com/product/arducam-cs-mount-lens-for-raspberry-pi-hq-camera-8mm-focal-length-with-manual-focus-and-adjustable-aperture/
32 GB SD card (5 units / pack) | off-the-shelf | sd_card | 1 | 5.6 | A micro sd card. 32GB should be enough, unless you retrieve the data from your Sticky Pi very infrequently (e.g. every other month). You need to [manually flash it](#_burn_sd_card) before inserting it into the [Raspberry Pi](#pi_zero) |  | amazon.ca | https://www.amazon.ca/KOOTION-Micro-10pcs-UHS-I-Memory/dp/B07XKX2NXP/ref=sr_1_10?keywords=micro+sd+16&qid=1573071363&sr=8-10
15cm camera flex cable ( (HBV RPI 160 FPC) | off-the-shelf | flex_cable | 1 | 1 | A flexible cable to connect that connects the [camera](#camera) to the [Raspberry Pi](#pi_zero) |  | aliexpress.com | https://www.aliexpress.com/item/32947418836.html?spm=a2g0o.cart.0.0.5eaa3c001hHS8H&mp=1
Microscope slide (standard: 75 x 25 mm) | off-the-shelf | glass_slide | 1 | 0.25 | A microscope slide that you can cut manually to make a lens cover on the [front of the camera box](#cambox_front) |  | amazon.ca | https://www.amazon.ca/AmScope-BS-50P-Microscope-Ground-Pre-Cleand/dp/B00NYI4RSG/ref=sr_1_9?crid=3BS2D29E3I3PE&keywords=microscope+slides+25+x+75&qid=1654125751&sprefix=microscope+slides+25+x+75+%2Caps%2C128&sr=8-9
Zip tie 12in | off-the-shelf | zip_tie_long | 2 | 0.1 | A pair of long zip ties to attach the [temperature probe](#dht22) |  | amazon.ca | https://www.amazon.ca/gp/product/B08N4B3PWC/ref=ppx_yo_dt_b_asin_title_o08_s00?ie=UTF8&psc=1
Zip tie 4in | off-the-shelf | zip_tie_short | 1 | 0.1 | A small zip tie to secure the cable during the [lightbox assembly](#_build_lightbox) |  | amazon.ca | https://www.amazon.ca/gp/product/B09TQXR4LR/ref=ppx_yo_dt_b_asin_title_o01_s00?ie=UTF8&psc=1
Male PCB headers (2x20, 2.54mm) | off-the-shelf | header_male_2x20 | 1 | 0.5 | PCB headers to solder on the [Raspberry Pi](#pi_zero) | You might also be able to buy raspberry pis with presoldered headers | aliexpress.com | https://www.aliexpress.com/item/32878530745.html?spm=a2g0o.productlist.0.0.69656a17yX6gny&algo_pvid=22c123fd-bad6-407e-8ab4-8394b37d1f72&algo_exp_id=22c123fd-bad6-407e-8ab4-8394b37d1f72-8&pdp_ext_f=%7B%22sku_id%22%3A%2265640776223%22%7D&pdp_pi=-1%3B2.74%3B-1%3B-1%40salePrice%3BCAD%3Bsearch-mainSearch
Male PCB headers (2x1, 2.54mm) | off-the-shelf | header_male_2x1 | 1 | 0 | PCB headers to solder on the testing slot of our [Raspberry Pi Hat](#pi_hat) | You can buy extra [2x20 headers](#header_male_2x20) and break them to 2x1. | aliexpress.com | https://www.aliexpress.com/item/32878530745.html?spm=a2g0o.productlist.0.0.69656a17yX6gny&algo_pvid=22c123fd-bad6-407e-8ab4-8394b37d1f72&algo_exp_id=22c123fd-bad6-407e-8ab4-8394b37d1f72-8&pdp_ext_f=%7B%22sku_id%22%3A%2265640776223%22%7D&pdp_pi=-1%3B2.74%3B-1%3B-1%40salePrice%3BCAD%3Bsearch-mainSearch
11pin 2.54mm PCB Screw Terminal Blocks | off-the-shelf | terminal_block | 1 | 1 | Small terminal blocks to solder on the back of our [Raspberry Pi Hat](#pi_hat) |  | aliexpress.com | https://www.aliexpress.com/item/1005002453862004.html?spm=a2g0o.9042311.0.0.7ffa4c4dxnouy5
dht22 temperature humidity sensor | off-the-shelf | dht22 | 1 | 4 | A humidity and temperature sensor |  | aliexpress.com | https://www.aliexpress.com/item/32901733917.html?spm=a2g0s.9042311.0.0.13564c4dqiiNwk
8mm push button | off-the-shelf | push_button | 1 | 1.4 | A waterproof stainless steel push button to insert in the [back of the camera box](#cambox_back) |  | aliexpress.com | https://www.aliexpress.com/item/4000043687021.html?spm=a2g0s.9042311.0.0.10444c4dmVJAuf
5mm orange led (WP7113SECK) | off-the-shelf | led_orange | 1 | 0.5 | A bright orange LED to insert in the [back of the camera box](#cambox_back) |  | DigiKey | https://www.digikey.com/en/products/detail/kingbright/WP7113SECK-J3/4753660
M12x1.2 cable gland | off-the-shelf | cable_gland | 1 | 1 | A stainless steel cable gland to insert in the [back of the camera box](#cambox_back) |  | aliexpress.com | https://www.aliexpress.com/item/1005001952576665.html?spm=a2g0o.cart.0.0.f9233c00dhXaEx&mp=1
CR1220 Coin Cell | off-the-shelf | coin_cell | 1 | 0.5 | A small battery for our [Raspberry Pi Hat](#pi_hat). This is necessary for the clock to keep accurate time. The battery should last for several years (though we are currently testing this). |  | amazon.ca | https://www.amazon.ca/Murata-CR1220-Battery-Lithium-Coin/dp/B0857H2TT7/ref=sr_1_5?crid=1E8KBJ9OX5NBA&keywords=coin+cell+CR1220&qid=1641554977&sprefix=coin+cell+cr1220%2Caps%2C223&sr=8-5
Silica desiccant bag | off-the-shelf | silica_bag | 2 | 0 | Silica bags to attach to the [Raspberry Pi sledge](#pi_sledge) to absorb moisture. |  | aliexpress.com | https://www.aliexpress.com/item/32772734661.html?spm=a2g0s.9042311.0.0.4cf04c4dv3gVAu
M3 x 20 self-tapping philips screws | off-the-shelf | screw_m3x20 | 4 | 0.1 | Screws to mount the [Raspberry Pi Zero](#pi_zero) and our [custom Raspberry Pi hat](#pi_hat) together on the [Raspberry Pi sledge](#pi_sledge). |  | aliexpress.com | https://www.aliexpress.com/item/32835249403.html?spm=a2g0s.9042311.0.0.27424c4dLTOcNH
M3 x 14 self-tapping philips screws | off-the-shelf | screw_m3x14 | 8 | 0.2 | Screws to lock the [back](#cambox_back) and [front](#cambox_front) of the camera box. |  | aliexpress.com | https://www.aliexpress.com/item/32835249403.html?spm=a2g0s.9042311.0.0.27424c4dLTOcNH
M3 x 6 self-tapping philips screws | off-the-shelf | screw_m3x6 | 10 | 0.1 | Screws to attach the [camera](#camera) to the [Raspberry Pi sledge](#pi_sledge), the [Raspberry Pi sledge](#pi_sledge) to the [back of the camera box](#cambox_back), and to mount the [DHT22](#dht22) in [its shell](#dht_shell) |  | aliexpress.com | https://www.aliexpress.com/item/32835249403.html?spm=a2g0s.9042311.0.0.27424c4dLTOcNH
M3 Stainless Acorn Nut Cap | off-the-shelf | nut_acorn_m3 | 8 | 0.4 | A set of acorn nuts to lock the [rods](#rod_m3x250) |  | aliexpress.com | https://www.aliexpress.com/item/32993886063.html?spm=a2g0o.cart.0.0.26443c00UEA2H1&mp=1
M3 x 250 threaded rods | off-the-shelf | rod_m3x250 | 4 | 5 | Four rods to serve as the "legs" that connect the lightbox to the camera box |  | aliexpress.com | https://www.aliexpress.com/item/4000990108860.html?spm=a2g0o.cart.0.0.79ae3c00cEblbO&mp=1
M3 nuts | off-the-shelf | nut_standard_m3 | 8 | 0 | A set of nuts to lock the [rods](#rod_m3x250) |  | aliexpress.com | https://www.aliexpress.com/item/32886183839.html?spm=a2g0o.productlist.0.0.20492e91xidtyr&algo_pvid=de29cb22-5ca3-4ce6-bd8d-55a5cf42d14e&algo_exp_id=de29cb22-5ca3-4ce6-bd8d-55a5cf42d14e-0&pdp_ext_f=%7B%22sku_id%22%3A%2265651913269%22%7D&pdp_pi=-1%3B1.43%3B-1%3B-1%40salePrice%3BCAD%3Bsearch-mainSearch
2-pin automotive connector male | off-the-shelf | auto_connector | 1 | 0 | A weather-proof power connector | Buy male-female pairs and use the female for the battery | aliexpress.com | https://www.aliexpress.com/item/32801400952.html?spm=a2g0o.order_list.0.0.43281802B3HZ5R
Custom Pi hat | custom-ordered | pi_hat | 1 | 40 | Our custom Raspberry Pi hat. This board is factory-assembled. All the necessary files and schematics are available on our [GitHub repository](https://github.com/sticky-pi/sticky-pi-device/tree/main/hardware/sticky_pi_hat) |  | PCBWay
Camera front | 3d-printed | cambox_front | 1 | 0 | The front of the camera box |  | doc.sticky-pi.com | https://github.com/sticky-pi/sticky-pi-device/blob/main/hardware/CAD/cambox_frontstl
Camera back | 3d-printed | cambox_back | 1 | 0 | The back of the camera box |  | doc.sticky-pi.com | https://github.com/sticky-pi/sticky-pi-device/blob/main/hardware/CAD/cambox_backstl
Camera tube | 3d-printed | cambox_tube | 1 | 0 | The core of the camera box |  | doc.sticky-pi.com | https://github.com/sticky-pi/sticky-pi-device/blob/main/hardware/CAD/cambox_tubestl
Camera visor | 3d-printed | cambox_visor | 1 | 0 | The visor to help shelter the lens cover from light and water |  | doc.sticky-pi.com | https://github.com/sticky-pi/sticky-pi-device/blob/main/hardware/CAD/cambox_visorstl
Raspberry pi sledge | 3d-printed | pi_sledge | 1 | 0 | The frame on which the electronics is mounted. It is housed in the camera box |  | doc.sticky-pi.com | https://github.com/sticky-pi/sticky-pi-device/blob/main/hardware/CAD/pi_sledgestl
Standoff | 3d-printed | pi_standoff | 4 | 0 | Standoffs to securely attach the [Raspberry Pi Zero](#pi_zero) to the [custom board](#pi_hat) |  | doc.sticky-pi.com | https://github.com/sticky-pi/sticky-pi-device/blob/main/hardware/CAD/pi_standoffstl
DHT22 shell | 3d-printed | dht_shell | 1 | 0 | An open enclosure to mount the [temperature and humidity sensor](#dht22) |  | doc.sticky-pi.com | https://github.com/sticky-pi/sticky-pi-device/blob/main/hardware/CAD/dht_shellstl
Light box front | 3d-printed | lightbox_front | 1 | 0 | The front of the lightbox. It acts as a diffuser | It must be made of a transparent material | doc.sticky-pi.com | https://github.com/sticky-pi/sticky-pi-device/blob/main/hardware/CAD/lightbox_frontstl
Light box back | 3d-printed | lightbox_back | 1 | 0 | The back of the lightbox. It holds the LED light strips |  | doc.sticky-pi.com | https://github.com/sticky-pi/sticky-pi-device/blob/main/hardware/CAD/lightbox_backstl
Drawer hook | 3d-printed | drawer_hook | 8 | 0 | A custom washer with a feature to retain the sticky trap drawer |  | doc.sticky-pi.com | https://github.com/sticky-pi/sticky-pi-device/blob/main/hardware/CAD/drawer_hookstl

## derived_parts
part | type | tag | description | note
Prepped Raspberry Pi Zero | derived | prepped_pi_zero | The [Raspberry Pi Zero](#pi_zero) with soldered [headers](#header_male_2x20) and the [flashed SD](#burnt_sd_card) card inside
Testable Sticky Pi | derived | testable_device | The Sticky Pi device, ready for quality control
Prepped Sticky Pi Hat | derived | prepped_hat | Our [custom Raspberry Pi hat](#pi_hat) with the [coin cell](#coin_cell) and the manually-soldered [pins](#header_male_2x1) and [terminal blocks](#terminal_block)
Prepped male power connector | derived | prepped_connector | The elongated [male automotive connector](#auto_connector)
Slab | derived | slab | The electronic components mounted on their frame
Mounted camera | derived | mounted_camera | The [camera](#camera) mounted on custom [the sledge](#pi_sledge)
Assembled camera box | derived | assembled_cambox | The fully assembled and tested camera box
Prepared lightbox back | derived | prepped_lightbox_back | The back panel of the light box with the [LED strips](#led_strips) embedded in [epoxy](#epoxy)
Assembled sledge | derived | assembled_sledge | The assembled sledge, a frame holding the [camera](#camera), [Raspberry Pi](#pi_zero) and [custom Pi hat](#pi_hat)
Epoxied camera box tube | derived | epoxied_tube | The [camera box tube](#cambox_tube) coated with [epoxy](#epoxy) and cured
Tested Sticky Pi | derived | tested_device | The Sticky Pi after quality control
Prepared camera box front | derived | prepped_cambox_front | The front of the camera box, with its sealed [glass window](#cut_slide)
Prepared DHT22 | derived | prepped_dht | The prepared [temperature and humidity sensor](#dht22), mounted on its [shell](#dht_shell)
Assembled light box | derived | assembled_lightbox | The complete and functional light box
Prepared camera box back | derived | prepped_cambox_back | The back of the camera box, with its [push button](#push_button), [cable gland](#cable_gland) and [LED](#led_orange)
Sticky Pi | derived | sticky_pi | The Sticky Pi, ready to use! Congratulations!
Flashed SD card | derived | burnt_sd_card | The [sd card](#sd_card) with our custom OS flashed on
Mounted camera box tube | derived | mounted_cambox_tube | The camera box tube with its front mounted on

## consumables
part | type | tag | quantity | price_per_device_CAD | description | note | provider | link
Epoxy | consumables | epoxy | 5ml | 0.005 | To seal, weatherproof and strengthen components |  | amazon.ca | https://www.amazon.ca/gp/product/B011T3CWPC/ref=ppx_yo_dt_b_asin_title_o02_s00?ie=UTF8&psc=1
Rainx | consumables | rainx | 1 spray | 0 | To make the [camera cover](#prepped_cambox_front) water repellent | Optional | amazon.ca | https://www.amazon.ca/gp/product/B000BVRZ74/ref=ppx_yo_dt_b_asin_title_o01_s00?ie=UTF8&psc=1
Liquid tape | consumables | liquid_tape | 1ml | 0 | To electrically insulate electrical connections |  | amazon.ca | https://www.amazon.ca/Gardner-Bender-LTB-400-4-Ounce-Electrical/dp/B000FPAN2K/ref=sr_1_2?keywords=liquid+tape&qid=1589929357&s=hi&sr=1-2
100% Silicone, transparent | consumables | silicone | 10ml | 0 | To seal the [lightbox](#assembled_lightbox) |  | amazon.ca | https://www.amazon.ca/gp/product/B000LGCEES/ref=ppx_yo_dt_b_asin_title_o05_s00?ie=UTF8&psc=1
Closed-cell insulating foam tape | consumables | foam | 15cm | 0 | To seal the [camera box](#assembled_cambox) |  | amazon.ca | https://www.amazon.ca/M-D-Building-Products-2758-Density/dp/B000CSIBCE/ref=sr_1_3?crid=2GG79BLS1NQO8&keywords=climaloc+foam&qid=1654124770&sprefix=climaloc+foam%2Caps%2C125&sr=8-3
Copper 30cm (24 AWG) | consumables | wire_copper | 24cm | 0 | To connect the [LED strips](#led_strips) inside the [lightbox](#assembled_lightbox) |  | amazon.ca | https://www.amazon.ca/gp/product/B00E4WL2D4/ref=ppx_yo_dt_b_asin_title_o04_s00?ie=UTF8&psc=1
Ribbon cables | consumables | cable_ribbon | 7x15cm | 0 | To connect peripherals to our [Raspberry Pi hat](#custom_pi_hat) |  | aliexpress.com | https://www.aliexpress.com/item/1005003090897764.html?spm=a2g0o.detail.1000060.1.3db8542ahBQ2Ut&gps-id=pcDetailBottomMoreThisSeller&scm=1007.13339.169870.0&scm_id=1007.13339.169870.0&scm-url=1007.13339.169870.0&pvid=e491e051-a6d5-469f-b0d5-a677642ccd8a&_t=gps-id:pcDetailBottomMoreThisSeller,scm-url:1007.13339.169870.0,pvid:e491e051-a6d5-469f-b0d5-a677642ccd8a,tpp_buckets:668%232846%238116%232002&&pdp_ext_f=%7B%22sceneId%22:%223339%22,%22sku_id%22:%2212000024023383531%22%7D
22AWG 2core cables | consumables | cable_twocore | 1m | 1 | To power the device and the [lightbox](#assembled_lightbox) |  | amazon.ca | https://www.amazon.ca/gp/product/B078TR567C/ref=ppx_yo_dt_b_asin_title_o02_s00?ie=UTF8&th=1
5V white LED | consumables | led_strips | 1.2m | 2 | To illuminate the [lightbox](#assembled_lightbox) |  | aliexpress.com | https://www.aliexpress.com/item/4000074675463.html?spm=a2g0o.order_list.0.0.21ef1802VsAzYQ
Heat-shrink tubing 1mm | consumables | heatshrink_small | 10cm | 0 | To insulate the pins of small components |  | aliexpress.com | https://www.aliexpress.com/item/1005002395602372.html?spm=a2g0o.order_detail.0.0.7130f19cJAoJ6N
Heat-shrink tubing 3mm red | consumables | heatshrink_red | 10cm | 0 | To splice and insulate the [2-core cables](#cable_twocore) |  | aliexpress.com | https://www.aliexpress.com/item/1005002395602372.html?spm=a2g0o.order_detail.0.0.7130f19cJAoJ6N
Heat-shrink tubing 3mm black | consumables | heatshrink_black | 10cm | 0 | To splice and insulate the [2-core cables](#cable_twocore) |  | aliexpress.com | https://www.aliexpress.com/item/1005002395602372.html?spm=a2g0o.order_detail.0.0.7130f19cJAoJ6N
Masking tape | consumables | tape_masking | 5cm | 0 | To help with the [lightbox assembly](#_assemble_lightbox) |  | amazon.ca | https://www.amazon.ca/Scotch%C2%AE-General-Purpose-Masking-Tape/dp/B00004Z47H/ref=sr_1_7?keywords=masking+tape&qid=1654126965&sr=8-7
Parchment paper | consumables | parchment_paper | 50cm | 0 | To prevent epoxy from sticking to worktops |  | amazon.ca | https://www.amazon.ca/Reynolds-Kitchens-Parchment-Premium-Non-Stick/dp/B0166SPEL4/ref=sr_1_6?keywords=parchment+paper&qid=1654127009&sprefix=parchm%2Caps%2C136&sr=8-6
Gloves | consumables | gloves | 2 | 0 | To work with epoxy
Duct tape | consumables | tape_duct | 10cm | 0 | To secure the [desiccant bags](#silica_bag) on the [pi sledge](#assembled_sledge)
Hot glue | consumables | hot_glue | 2g | 0 | To build the [camera cover](#prepped_cambox_front)
Modeling clay | consumables | clay | 5g | 0 | To build the [camera cover](#prepped_cambox_front) |  | amazon.ca | https://www.amazon.ca/Crayola-Modeling-Supplies-Stocking-Stuffers/dp/B011Q25T6K/ref=sr_1_15?crid=2U4AY2P90TOQ9&keywords=plasticine+modeling+clay+non+hardening&qid=1654127251&sprefix=plasticine%2Caps%2C124&sr=8-15
Lead-free solder | consumables | solder | ? | 0 | To assemble various electronics

## processes
name | tag | description | asset | note
Flash SD card | _burn_sd_card | 1. Install the [etcher](https://www.balena.io/etcher/) free software 2. Insert the SD card in the [microSD card adaptor](#sd_card_adaptor) 3. Select "Flash from URL" and enter {SD_IMAGE_LINK}. If you flash multiple SD cards, you can also download this file 4. Ensure you select the right SD card as the "target" and press flash 5. Verify that the operation was successful | https://figshare.com/ndownloader/files/36001211/preview/36001211/video_preview.mp4
Mount camera | _mount_camera | 1. Unbox the [Pi Camera HQ](#camera) and the [lens](#lens) 2. Remove the flex connector, tripod mount (using an [allen key](#-allen_key)), the C-CS adaptor, and the dust cap 3. Screw the lens tight on the camera (you should feel no resistance) 4. Use four [6mm m3 self-tapping screws](#screw_m3x6) to secure the camera on the [sledge](#pi_sledge). The connector of the camera should point towards the solid vertical wall of the sledge 5. Connect and thread the [flex connector](#flex_connector) as shown in the video | https://figshare.com/ndownloader/files/36001229/preview/36001229/video_preview.mp4 | Be gentle when handling optics and electronics. You might need to file the screw holes of the camera a little bit, as in the video, depending on the screws you got
Prepare Pi Zero | _prep_pi_zero | 1. Unbox the [Raspberry Pi Zero](#pi_zero) 2. Solder the [male headers](#header_male_2x20) 3. Insert the [flashed SD card](#burnt_sd_card) | https://figshare.com/ndownloader/files/36001253/preview/36001253/video_preview.mp4 | Whilst soldering, you want to sink the heat (for instance, using a [breadboard](#-breadboard)) and go quickly so not to stress the board
Prepare Pi hat | _prep_pi_hat | 1. Solder the [terminal block](#terminal_block) onto our [custom Pi hat](#pi_hat). The holes should face towards the centre of the board 2. Solder a [2x1 male header](#header_male_2x1) on the "testing" slot of our [custom Pi hat](#pi_hat) 3. Insert a [coin cell](#coin_cell) in the battery slot of our [custom Pi hat](#pi_hat) | https://figshare.com/ndownloader/files/36001250/preview/36001250/video_preview.mp4
Assemble sledge | _assemble_sledge | 1. Connect the [prepared Pi hat](#prep_pi_hat) to the [prepared Raspberry Pi Zero](#prepped_pi_zero) 2. Position the Pi Zero on the [sledge with the mounted camera](#mounted_camera) so that the Pi's camera connector face towards the camera 3. Connect the flex connector to the Raspberry Pi 4. Align the board corner holes with the holes on the sledge 5. Screw [self-tapping 20mm m3 screws](#screw_m3x20), one by one, through corner holes, from the top side whilst holding a [standoff](#pi_standoff) between the Pi hat and the Pi zero. 6. Use duct tape to secure two [silica desiccant bags](#silica_bag) on the sledge | https://figshare.com/ndownloader/files/36001208/preview/36001208/video_preview.mp4 | Ideally, you can dry your silica bag (e.g. in an oven) before assembly.
Epoxy tube | _epoxy_tube | 1. Lay some [parchment](#parchment_paper) on your bench 2. Prepare some [epoxy](#epoxy), following the manufacturer's instruction (10mL should be enough for a tube). 3. Use a paper towel to apply epoxy on the [camera box tube](#cambox_tube), both inside and outside 4. Remove excess, in particular on the edges 5. Let the part cure for 24h, sideways, on the parchment | https://figshare.com/ndownloader/files/36001214/preview/36001214/video_preview.mp4 | Rather than epoxying parts one by one, it is often less wasteful and more efficient to do it in a batch.
Prepare cambox front | _prep_cambox_front | 1. Use a [custom glass cutting gauge](#-glass_cut_gauge) and a [glass cutting pencil](#-glass-pencil) to cut a [microscope glass side](#glass_slide) into a **25x25mm** piece 2. Use a [hot glue gun](#-hot_glue_gun) and [hot glue](#hot_glue) to glue the cut slide to the **outside** of the [camera box front](#cambox_front) 3. When set, use [modeling clay](#clay), cut with a [square cutter](#square_cutter), to make a casting mask on the **inside** of the camera box front 4. Press the clay square on the glass, visually ensure it is sealed, and check there is, at least, a 3mm margin between the clay and the edge of the glass 5. Lay the camera box front on [parchment](#parchment_paper) 6. Cast [epoxy](#epoxy) around the glass to embed it in the camera box front 7. Wait 24h, remove the clay, and clean with detergent and a [toothbrush](#-tooth-brush) 8. Cut an 80x5mm strip of [foam](#foam) and insert it in the groove of the camera box front | https://figshare.com/ndownloader/files/36001235/preview/36001235/video_preview.mp4 | Rather than epoxying parts one by one, it is often less wasteful and more efficient to do it in a batch.
Mount cambox front | _mount_cambox_front | 1. Clean the [prepared camera box front](#prep_cambox_front) 2. Place the [prepared camera box front](#prep_cambox_front) on the [epoxied camera box tube](#epoxied_tube), so that the holes align. You can push and twist the camera box front on the tube. 3. Place the [camera box visor](#cambox_visor) on the camera box front, so that the pointy triangular edges align. 4. Use four [self-tapping 14mm m3 screws](#screw_m3x14) to lock all three parts together. Ensure you press the parts together to improve sealing | https://figshare.com/ndownloader/files/36001226/preview/36001226/video_preview.mp4 | If the screw holes of the camera box tube have some epoxy in them, you may want to tap them with a small drill bit
Prepare humidity sensor | _prep_dht | 1. Split a [ribbon cable](#ribbon cable) to keep three cores, and cut a 30cm long piece. Pick three distinctive colours 2. Peel both ends of the ribbon cables ([wire strippers](#-strippers) can be useful) 3. Slide 1cm long [narrow heatshrink](#heatshrink_small) onto the three cores, on one end of the ribbon cable. 4. Solder the three cores on the `+`, `data` and `-` pins of the [DHT22](#dht22) 5. Shrink the heatshrink around the solder 6. Thread the ribbon cable inside the hole of the [dht22 shell](#dht_shell) and screw the dht22 inside the shell with two [self-tapping 6mm m3 screws](#screw_m3x6) 7. Apply some liquid tape to fill the cable hole in the dht shell and on the connections of the DHT22, to prevent corrosion | https://figshare.com/ndownloader/files/36001241/preview/36001241/video_preview.mp4
Prepare cambox back | _prep_cambox_back | 1. Screw the [cable gland](#cable_gland) tight so that the long part is on the flat (outside) side of the [camera box back](#cambox_back). Use a [14mm wrench](#-wrench_14mm) 2. Screw the [push button](#push_button) from the outside. You might need [pliers](#-pliers_large) 3. Shorten the legs of the [orange LED](#led_orange) to about 5 and 6mm, for `-` and `+` pins. And insert the LED in the inside of the camera box back 4. Split, cut, and strip the ends of two pieces of two-core 10mm of [ribbon cable](#ribbon_cable). Choose sensible and consistent colours 5. Using [large heatshrink](#heatshrink_black), solder one end of the ribbon cable to push button. 6. Using [small heatshrink](#heatshrink_small), solder one end of the ribbon cable to the orange LED. Keep track of which colour is the `+` (long leg). Use [liquid tape](#liquid_tape), to seal the LED inside its socket. Allow for the liquid tape to cure (a few minutes) 7. Cast a single drop of [epoxy](#epoxy) on the outside hole of the LED socket. Allow 24h for the epoxy to cure 8. Cut a 180x5mm strip of [foam](#foam) and insert it in the groove of the camera box front | https://figshare.com/ndownloader/files/36001232/preview/36001232/video_preview.mp4 | Rather than epoxying parts one by one, it is often less wasteful and more efficient to do it in a batch.
Prepare connector | _prep_connector | 1. Cut and strip the ends of 10cm of a [two-core 24AWG](#cable_twocore) wire 2. Strip the end of the [male two-pin automotive connector](#auto_connector) 3. Using [black](#heatshrink_black) and [red heatshrink](#heatshrink_red), splice and solder cables together. Ensure none of the cable cores are exposed. | https://figshare.com/ndownloader/files/36001238/preview/36001238/video_preview.mp4
Merge sledge and prepped cambox back | _merge_sledge_back | 1. Align the board [prepared camera box back]() with the [assembled sledge](). The pointy edge of the camerabox back shows the top. As shown in the video, the camera connecto faces to the left, when looking in the same direction as the camera. 2. Screw four [self-tapping 6mm m3 screws](#screw_m3x20), one by one, through the corner holes whilst holding the parts togetherx` | https://figshare.com/ndownloader/files/36001223/preview/36001223/video_preview.mp4
Prepare lightbox back | _prep_lightbox_back | 1. Cut six pieces 16.5cm (i.e. 10 LEDs) of [white LED strips](#led_strips) 2. Use [wire clippers](#-wire-clippers) to remove all the `+` pads on one side of the six LED strip sections, and all the `-` on the other side 3. Place the LED strips inside the grooves in the lightbox back, so that all the terminal `+` pads are on one side and all the `-` on the other side 4. Cut two 12cm of [copper wire](#wire_copper). Use masking tape to keep the copper on top of the pads, and solder to join all `+` on one side and all `-` on the other 5. Cut 70cm of [two-core 24 AWG cable](#cable_twocore). Strip the ends 6. Insert the cable from the outside of the lightbox back, through the narrow round hole, and solder the red to a `+` pad, and the black to `-` a pad, on the first LED strip available 7. Use a [small zip tie](#zip_tie_small), on the inside of the box to lock the cable in place | https://figshare.com/ndownloader/files/36001232/preview/36001232/video_preview.mp4 | Remember to test the lightbox!
Assemble lightbox | _assemble_lightbox | 1. To test the lightbox, connect the external end of the 70cm cable to a 5V power supply (e.g. a stripped USB cable will do). All lights should brighten. 2. Cast 50mL of [epoxy](#epoxy) inside the lightbox to seal the lights and wires in place 3. Allow 24h to cure 4. Re-test the light box as above 5. Apply some [silicone](#silicone) in the groove of the [lightbox front](#lightbox_front) 6. Put the front and back of the lightbox together, press and allow for the silicone to cure | https://figshare.com/ndownloader/files/36001205/preview/36001205/video_preview.mp4
Connect cambox to peripherals | _merge_slab_dht_lightbox | 1. Through the cable gland, in the [back of the camera box](#slab), thread, the [assembled dht22 sensor](#prepped_dht), the [prepped power connector](#prepped_connector) and the cable from the [lightbox](#assembled_lightbox) 2. Use a [small screw driver](#-screwdriver_flat) and a pair of forceps to insert and screw the cable end one by one. 3. Each of the 11 terminal block sockets on our [custom Pi hat] is labelled. From left to right, connect the battery (`+`, `-`), the push button (2 slots), the orange LED (`+`, `-`), the DHT22 (`+`, data, `-`), the lightbox (i.e., flash: `+`, `-`). | https://figshare.com/ndownloader/files/36001220/preview/36001220/video_preview.mp4 | It is important to quality control this step. You want to try to remove every individual cable with forceps and ensure it holds. You should visually check that each cable is at the right place. Also, ensure no loose copper stand sticks out of any socket.
Test device | _test_device | Testing is is described in the [user manual section](#user-manual). Briefly: 1. Prepare the following accessories for testing: * a female header testing bridge * a female auto connector connected to a 5V supply (e.g. USB) * a testing rig, with a platform to hold the device and wall at focusing distance * an [HDMI cable](#-hdmi_to_micro) cable and a monitor * a data harvester ready to use 2. Connect the Raspberry Pi to the screen, with the HDMI cable 3. Connect the two testing pins of our custom Pi hat using the testing bridge 4. Power the board with the female auto connector. **Do not unplug**. You should see the orange LED turning on immediately 5. After a few seconds, the screen should display some information. The first time a new system is booted, it initialises itself (e.g. format the SD card). This should take a minute or so. The system will automatically reboot in testing mode afterwards (if the testing bridge is on). 6. The system will then go through a series of interactive checks to test and set the peripherals (sensors, LED, button, camera) 7. Once done, the system can be paired to a data harvester, using the mobile app
Assemble camera box | _assemble_cambox | 1. Insert the [tested device](#tested_device) in the [camera tube](#mounted_cambox_tube), the lens of the camera should fit in the slot on the [camera box front](#cambox_front) 2. Use four [self-tapping 14mm m3 screws](#screw_m3x14) to attach the [tested device](#tested_device) to the camera box tube. Ensure you press the parts together to improve sealing. 3. Use two [long zip ties](#zip_tie_long) to attach the [dht22 sensor and its shell](#prepped_dht) around the tube. The sensor should face down, away from the pointy edges of the camera box. 4. Use the [14mm wrench](#-wrench_14mm) to tighten the [cable gland](#cable_gland). | https://figshare.com/ndownloader/files/36001202/preview/36001202/video_preview.mp4
Merge camera box to lightbox | _merge_cambox_lightbox | 1. Screw an [m3 nut](#nut_standard_m3) in both ends of four [25cm m3 rods](#rod_m3x250), about 3cm along 2. Fit a [drawer hook](#drawer_hook) on each m3 rod, so that the flat side matches the m3 screw. Insert this end of the rods on the lightbox, from the front. Use an [acorn nut](#nut_acorn_m3), from the back of the lightbox to lock the rod in place. Just hand tighten for now 3. Coil the lightbox cable around the bottom left rod, to avoid having a loose cable. 4. Repeat the same process as 2., also using drawer hooks, as washers, for the [assembled camera box](#assembled_cambox). **Align the tops**: the pointy edges of the camera box, and the visor, align with the top of the light box. Check the pictures if unsure 5. Use a pair of [5.5mm Wrenches](#-wrenches) to first tighten the acorn nuts, then each m3 nut against its acorn nut | https://figshare.com/ndownloader/files/36001217/preview/36001217/video_preview.mp4

## tools
part | tag | number | description | price | provider | link
Flat screwdriver | -screwdriver_flat
Phillips #1 screwdriver | -screwdriver_philips | 1
power drill | drill | 1
3mm drill bit | drill_bit | 1
wire strippers | -strippers | 1
wire clippers | -clippers | 1
14mm wrench | -wrench_14mm
fine needle-nosed pliers | -pliers | 1
large pliers | -pliers-large | 1
curved forceps | forceps | 1
soldering iron | soldering_iron | 1
heat gun | heat_gun | 1
scissors | scissors | 1
toothbrush | toothbrush | 1
square cutter | square_cutter | 1 |  |  | amazon.ca | https://www.amazon.ca/Sculpey-AMM1016-Premo-Cutter-Pkg-Square/dp/B07STG9QNT/ref=sr_1_9?crid=1MVST1VB944YG&keywords=square+cutters+for+modeling+clay&qid=1654127612&sprefix=square+cutters+for+modeling+cla%2Caps%2C145&sr=8-9
hot glue gun | -hot_glue_gun | 1
Glass cutting pencil | -glass-pencil |  |  |  |  | https://www.amazon.ca/gp/product/B0834S18M3/ref=ppx_yo_dt_b_search_asin_title?ie=UTF8&psc=1
custom glass cutting gauge | -glass_cut_gauge | 1
5.5mm Wrenches | wrenches | 2
Micro Sd card adaptor | sd_card_adaptor | 1
1.5mm allen key | -allen_key |  | To remove the tripod mount on the [Pi Camera HQ](#camera)
HDMI cable | -hdmi_cable
HDMI Monitor | -monitor
breadboard | -breadboard
Syringe 10ml | -syringe

## accessories
part | type | tag | quantity | price_per_device_CAD | description | provider | link
Solar panel charger | accessories |  |  |  |  | aliexpress.com | https://www.aliexpress.com/item/33034581059.html?spm=a2g0o.order_list.0.0.21ef1802VsAzYQ
3.7V, 5Ah tablet battery | off-the-shelf |  | 1 | 12.6 |  | aliexpress.com | https://www.aliexpress.com/item/32791077730.html?spm=a2g0s.9042311.0.0.570e4c4dLTE9Zd
Card drawer | 3d-printed | card_drawer | 2 |  |  |  | TODO link to github stl

## assets_to_make
tag | type | image | description
pi_zero | primary_part
camera | primary_part
lens | primary_part
sd_card | primary_part
flex_cable | primary_part
glass_slide | primary_part
zip_tie_long | primary_part
zip_tie_short | primary_part
header_male_2x20 | primary_part
header_male_2x1 | primary_part
terminal_block | primary_part
dht22 | primary_part
push_button | primary_part
led_orange | primary_part
cable_gland | primary_part
coin_cell | primary_part
silica_bag | primary_part
screw_m3x20 | primary_part
screw_m3x14 | primary_part
screw_m3x6 | primary_part
nut_acorn_m3 | primary_part
rod_m3x250 | primary_part
nut_standard_m3 | primary_part
auto_connector | primary_part
pi_hat | primary_part
cambox_front | primary_part
cambox_back | primary_part
cambox_tube | primary_part
cambox_visor | primary_part
pi_sledge | primary_part
pi_standoff | primary_part
dht_shell | primary_part
lightbox_front | primary_part
lightbox_back | primary_part
drawer_hook | primary_part
prepped_pi_zero | derived_part
testable_device | derived_part
prepped_hat | derived_part
prepped_connector | derived_part
slab | derived_part
assembled_cambox | derived_part
prepped_lightbox_back | derived_part
assembled_sledge | derived_part
epoxied_tube | derived_part
tested_device | derived_part
prepped_cambox_front | derived_part
sledge | derived_part
prepped_dht | derived_part
assembled_lightbox | derived_part
prepped_cambox_back | derived_part
sticky_pi | derived_part
cut_slide | derived_part
burnt_sd_card | derived_part
epoxy | consumable
rainx | consumable
liquid_tape | consumable
silicone | consumable
foam | consumable
wire_copper | consumable
cable_ribbon | consumable
cable_twocore | consumable
led_strips | consumable
heatshrink_small | consumable
heatshrink_red | consumable
heatshrink_black | consumable
tape_masking | consumable
parchment_paper | consumable
gloves | consumable
tape_duct | consumable
hot_glue | consumable
clay | consumable
screwdriver | tool
drill | tool
drill_bit | tool
strippers | tool
clippers | tool
pliers | tool
forceps | tool
soldering_iron | tool
heat_gun | tool
scissors | tool
toothbrush | tool
square_cutter | tool
hot_glue_gun | tool
glass_cut_gauge | tool
wrenches | tool
sd_card_adaptor | tool